|  |
| --- |
| Comfort probe      Application information |
|  |

Pos: 1 /TD/Kurzanleitungen/testo 480/Behaglichkeitssonde 0628 0143 @ 9\mod\_1309359060255\_79.docx @ 81219 @ 555 @ 4

##### Application

The comfort probe 0628 0143 is used in conjunction with testo 480 to determine the air velocity and draught risk in accordance with DIN EN 13779.

|  |  |
| --- | --- |
| Info | Observe the information on the measuring process in the instrument instruction manual. |

##### Overview

|  |
| --- |
| 0628 0143 |
|  |

1 Sensors

|  |
| --- |
|  |
| CAUTION |
| **Damage to sensor!**  > Do not touch the sensor.  > Cover over with the protective sleeve after the measurement. |
|  |

2 Handle

3 Flow direction (flow against testo logo)

4 Connection for plug-in head cable (art. no. 0430 0100)

##### Technical data

| Feature | Values |
| --- | --- |
| Flow measuring range | 0 to +5 m/s |
| Temperature measuring range | 0 to 50 °C (probe head) |
| Operating range, handle | 0 to 40 °C |
| Accuracy (at nominal temperature 22 °C) ± 1 digit | ±0.5 °C  ±(0.03 m/s + 4% of meas. val.)   |  |  | | --- | --- | |  |  | | Info | At low air velocities, slightly increased temperature display! | |  |  | |
| Absolute pressure sensor | measuring range: +700...+1100 hPa  Accuracy: ±3 hPa |
| Adjustment conditions | Wind tunnel, adjusted in free jet Ø 350 mm, reference pressure 1013 hPa, flow against testo logo |

|  |  |
| --- | --- |
| Info | The digital probe allows measurement values to be processed directly in the probe. This technology eliminates instrument measurement uncertainty.  For calibration, the probe alone (without the hand instrument) can be sent away.  Entering the determined adjustment data via the EasyClimate software generates a zero-error display. |

=== Ende der Liste für Textmarke Inhalt ===

|  |
| --- |
|  |
|  |
|  |
|  |
|  |
|  |
|  |